## Master
| MASTER CHECKLIST FOR GLOBAL INITIAL/FINAL PROCEDURES | Day:
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
AT THE BEGINNING OF THE EXPERIMENTAL RUN (ONE OR MULTIPLE SHOTS)
E1 | Ensure gas supplies (for the Driver and the Test Gases) are OFF
E2 | Ensure all mechanical operators to wear proper PPE
E3 | Ensure every safety systems are ON and OK (fire alarm, emergency lighting, gas detection…)
E4 | Ensure UPS of control room and experimental hall are ON, supplying computers, control panel, gas detection system, and instrument racks. Ensure that the PLC control systems (in the control room and the experimental hall rack) are ON
E5 | All doors to remain unlocked as normal for fire escape/access
R1 | Turn ON necessary electrical equipment (computers, oscilloscopes, etc…)
FOR EACH SHOT
R2 | Make a realistic assessment on whether the test can be completed within the imparted time (e. g. if it is 16:00 maybe it is not a good idea to start a test if you plan to leave at 17:00).
R3 + E6 | Agree and sign the global shot setup parameters sheet (driver, compression tube, and test section gas fillings, vacuum parameters, diagnostics, etc…). Ensure parameters are nominal for shot and all equipment is in good condition. Two signatures are needed in the compliance form.
R4 + E7 | Ensure that at least one researcher and the chief engineer are present during the test, one to remain in control room throughout.
E8 | Undertake CC cleaning (if necessary) and CT and ST cleaning. Perform hardware inspection. Place diaphragm of appropriate sizes in the CC-CT, CT-ST, and ST-DT sections for the planned shot.
E9 | Ensure there is no leaks of hydraulic fluid. Undertake hydraulic closure procedure for the CC-CT, CT-ST, and ST-DT sections (See Procedimento HPL-012020).
R5 | Initiate Vacuum Procedure.
R6 | After the necessary time (experiment-dependent), initiate Combustion Driver procedure.
R7 | After Vacuum procedure concludes, initiate Test Gas Filling procedure.
R8 | After shot is concluded, ensure that all the relevant data (pressures, diagnostics, etc…) are recorded, and that pressure inside the chamber (measured in the Dump Tank station is around atmospheric pressure, accounting for the sensor noise.
R9 | Allow for additional 5min wait time for safety reasons.
R10 | Turn ON the water cooling circuit pump. Turn ON the ST primary pump to vent out any leftover Helium and Water Vapour. When pressure is lower than 100mbar, turn OFF the ST primary pump. Turn OFF the water cooling circuit pump.
E10 | Undertake hydraulic opening procedure (See Procedimento HPL-012020).
E11 | Purge condensed water from the Dump Tank.
E12 | Remove the burst-ed CC-CT, CT-ST, and ST-DT diaphragms.
E13 | Check mechanical parts for damage, record for future maintenance.
R11 + E14 | Sign test report with this form attached. End of test.
AT THE END OF THE EXPERIMENTAL RUN
R12 | Turn OFF relevant electrical equipment (computers, oscilloscopes, etc…).
E15 | General cross-check of the experimental hall, control room, and compressor room for any anomalies.
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Combustion Driver
| COMBUSTION CHAMBER FILLING AND FIRING | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
E2 | Disable electrical supplies to all non-ATEX electrical equipment in the ATEX area
E3 | Ensure all doors and gate from experimental hall are CLOSED
E4 | Place the recoil dampening bags (Aluminium bullet bags) in contact with the rear of the combustion chamber. Ensure these do not block the laser optical path.
E5 | Turn ON laser ignition system
E6 | Check for lack of human presence in the gas compressor room and CLOSE door
E7 | Leave experimental hall and make visual check and audible warning
E8 | CLOSE control room/lab door
E9 | OPEN gas storage doors; Turn ON gas supplies; CLOSE and LOCK gas storage doors
R1 | Log in to control console
R2 | Enter normalized parameters from log sheet
R3 | Turn on Kistler & start archive engine
R4 | Initiate filling sequence
R5 | Record the partial filling pressures
R6 | Check that fill is nominal and that the correct final pressure has been reached
R7 | Check that CC is properly insulated from the piping system (Maximator valves closed)
R8 | Ensure that the Vacuum and test gas filling procedures have concluded successfully and that the shock-detection and diagnostic systems are ready for recording the shot and waiting for the triggers recorded at shock wave passage
R9 | Prepare the Kistler gauge system to record the ignition sequence; program all the necessary triggers (oscilloscopes, acquisition cards, etc...), start the laser flash lamp
R10 | Fire the laser ignition system
R11 | Record full pressures and volumes on log
R12 | Set control system to idle
E10 | Open gas storage doors; Turn OFF gas supplies; Close and lock gas storage doors
E11 | Enter lab and open doors. Perform a visual inspection
E12 | Turn OFF laser ignition system and remove the recoil dampening bags.
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Vacuum
| VACUUM PROCEDURE | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
R1 | Verify that all the sections of the shock-tube (CC-CT, CT-ST, ST-DT) are CLOSED with a diaphragm inside, and pressurized to the preset values in each station. Report the hydraulic pressure in each section.
R2 | Turn ON the water cooling pump and check that the water is running inside the closed circuit
R3 | Turn ON the two turbomolecular pump fans
R4 | Turn ON the two turbomolecular pump controllers and ensure that the pumps are levitating nominally without any error messages
R5 | Ensure that the two ST gate valves are OPEN and that the two mechanical arms are in the OPEN position
R6 | Ensure that all the pressure sensors (1x CT, 2x ST, and 2x DT) are ON (turn then ON when required)
R7 | Ensure that the CT ball valve is in the pump ↔ CT section position and that the DT ball valve is in the pump ↔ DT section position
R8 | Ensure that all the control software is ON (boot it as needed) and that the archive engine is active
R9 | Turn ON the DT primary pump
R10 | After XXX seconds turn ON the CT primary pump
R11 | After XXX seconds turn ON the ST primary pump
R12 | After XXX seconds, ensure that the CT massflow controller is in the CLOSED position, then switch the CT ball valve to the gas-filling system ↔ CT section position, wait YYY seconds, and switch back to the pump ↔ CT section position, to purge the CT gas lines downstream of the massflow controller
R13 | After pressure in the ST section falls in the 1e-1 mbar range, start the acceleration of both ST turbomolecular pumps
R14 | When the required pressures in all the sections (CT, ST, DT) are met, CLOSE the two ST gate valves, switch the CT ball valve to the gas-filling system ↔ CT section position, and switch the DT ball valve to the purge line ↔ DT section position
R15 | After 5min, OPEN the two ST gate valves, and switch back the ball valves to the pump ↔ CT section and pump ↔ DT section positions
R16 | After 5min, CLOSE the two ST gate valves, and switch the ball valves to the gas-filling system ↔ CT section and vent line ↔ DT section positions
R17 | Give the OK signal for initiating test gas filling
R18 | Start the deceleration of both ST turbomolecular pumps
R19 | Turn OFF the CT and DT primary pumps
R20 | Once the ST turbomolecular pumps return to the levitation position (0 rpm), turn OFF the ST primary pump
R21 | Turn OFF the two turbomolecular pump controllers
R22 | Turn OFF the two turbomolecular pump fans
R23 | Turn OFF the water cooling circuit pump
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Test Gases (CT, ST)
| TEST GAS (CT, ST) FILLING PROCEDURE | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
R1 | Wait for the OK signal from the Vacuum system for initiating test gas filling
TODO
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Shock Detection System
| SHOCK DETECTION SYSTEM PROCEDURE | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
R1 | Wait for the OK signal from the Test Gases subsystem for initiating procedure
TODO
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Optical Diagnostics
| OPTICAL DIAGNOSTICS PROCEDURE | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
R1 | Wait for the OK signal from the Test Gases subsystem for initiating procedure
TODO
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge:

## Microwave Diagnostics
| MICROWAVE DIAGNOSTICS PROCEDURE | Test number: |  |  | S____
Date/Time Start
# | Test procedure | Check(a) signature |  |  | Notes
|  | V | ! | X
E1 | Verify that all other general and subsystem safety conditions are met (see Master checklist)
R1 | Wait for the OK signal from the Test Gases subsystem for initiating procedure
TODO
Date/Time Stop
(a) V – No problem ¦ ! – Attention required ¦ X – Urgent repair action
R - Researcher in charge: E - Chief Engineer in charge: